## Slide 1: Confocal Microscopy
Pablo Ariel, Ph.D.
Microscopy Services Laboratory
Department of Pathology and Laboratory Medicine
University of North Carolina, Chapel Hill
June 2020

## Slide 2
Why confocal ?
What problem does a confocal solve ?

## Slide 3
Sample
Sensitive area detector
Detection in parallel
Out of focus objects are blurred
Fluorophores
Focal plane
Lenses et al
This is what a confocal solves
Widefield fluorescence microscopy
Lamp
Notes: Highly simplified ray diagrams
Ray tracing: https://ricktu288.github.io/ray-optics/

## Slide 4
Notes: https://patents.google.com/patent/US3013467A/en

## Slide 5
Example: 15um kidney section, 63X/1.4NA objective
How does a confocal do this?
Bonus: Compare for yourself at https://www.olympus-lifescience.com/en/microscope-resource/primer/java/confocalvswidefield/
Notes: Modified from https://doi.org/10.1016/B978-0-12-420138-5.00007-0

## Slide 6
Detection in parallel
Out of focus objects are blurred
Sample
Sensitive area detector
Lenses et al
(Less) sensitive spot detector
Focal plane
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser scanning confocal microscopy
Widefield fluorescence microscopy
Light from out of focus objects removed
Lamp
Laser

## Slide 7
Detection in parallel
Out of focus objects are blurred
Sample
Sensitive area detector
Lenses et al
(Less) sensitive spot detector
Focal plane
Light from out of focus objects removed
Detection in series (“scanning”)
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser scanning confocal microscopy
Widefield fluorescence microscopy
Lamp
Laser

## Slide 8
Out of focus objects are blurred
Light from out of focus objects removed
Detection in parallel
Sample
Sensitive area detector
Lenses et al
(Less) sensitive spot detector
Focal plane
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Detection in series (“scanning”)
Laser scanning confocal microscopy
Widefield fluorescence microscopy
Lamp
Laser

## Slide 9
Out of focus objects are blurred
Light from out of focus objects removed
Detection in parallel
Sample
Sensitive area detector
Lenses et al
(Less) sensitive spot detector
Focal plane
Detection in series (“scanning”)
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser scanning confocal microscopy
Widefield fluorescence microscopy
Lamp
Laser

## Slide 10
Laser scanning is 2D
X
Y
https://www.olympus-lifescience.com/en/microscope-resource/primer/java/galvanometerscanning/
Notes: Heavily modified from Figure 4: http://zeiss-campus.magnet.fsu.edu/articles/spinningdisk/introduction.html

## Slide 11
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser scanning confocal microscopy
Laser

## Slide 12
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser scanning confocal microscopy
Laser
Critical for removing out of focus light

## Slide 13
How do we set the pinhole size ?
~1 Airy unit is usually a good compromise
Notes: Graphics modified from
https://ieeexplore.ieee.org/document/8039195
https://warwick.ac.uk/fac/sci/lifesci/research/facilities/imaging/sp5booking/imaging-pinhole/
https://doi.org/10.1016/B978-0-12-420138-5.00007-0

## Slide 14
Confocals can image multiple fluorophores
3 PMTs
Pinhole
Dichroics and filters
Dichroic
Excitation laser
Scanning
mirrors
Objective and sample
Notes: http://cshprotocols.cshlp.org/content/2014/10/pdb.top071795

## Slide 15
Modern confocals have flexible wavelength detection windows
Example: Zeiss LSM710 light path
(two of these in cores on campus)
Video: Zeiss LSM980 - https://zeiss.wistia.com/medias/4ehx150o5j

## Slide 16
Different companies do this in different ways
Zeiss LSM710 – Diffraction grating
Olympus FV3000 – Volume phase hologram
Leica SP8 - Prism
Nikon A1 – many options
Notes: https://applications.zeiss.com/C125792900358A3F/0/55FE5FC0794DD7E0C12583C10039E6A3/$FILE/EN_41_011_082_LSM980_rel1-1.pdf
https://www.olympus-lifescience.com/en/.downloads/download/?file=855640653&fl=en_US&inline

## Slide 17
Laser scanning confocals can do brightfield at the same time
This detector gets brightfield image at same time as fluorescence
Leica SP8

## Slide 18
X
Y
Confocal allows 3D imaging
Recipe:
Scan laser in 2D
Move sample in Z
Repeat
Z
Notes: Modified from Figure 4: http://zeiss-campus.magnet.fsu.edu/articles/spinningdisk/introduction.html

## Slide 19
Confocal allows 3D imaging
Recipe:
Scan laser in 2D
Move sample in Z
Repeat
Pollen grain serial optical sections with a confocal
Reconstruct 3D
Notes: Modified from https://www.olympus-lifescience.com/en/microscope-resource/primer/techniques/confocal/confocalintro/

## Slide 20
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Problems with laser scanning confocal
Laser

## Slide 21
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Not very sensitive
Problems with laser scanning confocal

## Slide 22
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Slow
Not very sensitive
Problems with laser scanning confocal

## Slide 23
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Slow
Damaging
Not very sensitive
Problems with laser scanning confocal

## Slide 24
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Not very sensitive
Increasing sensitivity of laser scanning confocal

## Slide 25
Standard photomultiplier tubes are noisy and not very sensitive
Increasing sensitivity of laser scanning confocal
Problem
Notes: Information from
Olympus: https://www.olympus-lifescience.com/en/laser-scanning/fv3000/#!cms[focus]=cmsContent6335
Leica: https://downloads.leica-microsystems.com/Leica%20HyD/Dedicated%20Articles/Photoncounting_Technological%20Readings_Jul2015.pdf
https://journals.plos.org/plosone/article?id=10.1371/journal.pone.0214659
Zeiss:

## Slide 26
Standard photomultiplier tubes are noisy and not very sensitive
Increasing sensitivity of laser scanning confocal
Problem
Solution
Better detectors ($$$)
GaAsP PMTs (many companies)
HyDs (PMT+APD, Leica)
Notes: Information from
Olympus: https://www.olympus-lifescience.com/en/laser-scanning/fv3000/#!cms[focus]=cmsContent6335
Leica: https://downloads.leica-microsystems.com/Leica%20HyD/Dedicated%20Articles/Photoncounting_Technological%20Readings_Jul2015.pdf
https://journals.plos.org/plosone/article?id=10.1371/journal.pone.0214659
Zeiss:

## Slide 27
Pinhole blocks light
Increasing sensitivity of laser scanning confocal
Problem
Image of 100nm beads with Zeiss 880 LSM
Pinhole
Notes: Information from:
https://warwick.ac.uk/fac/sci/lifesci/research/facilities/imaging/sp5booking/imaging-pinhole/

## Slide 28
Pinhole blocks light
Increasing sensitivity of laser scanning confocal
Problem
Solution
Airyscan (Zeiss)
Rethink pinhole and point detection ($$$)
Notes: Information from:
https://www.zeiss.com/microscopy/us/products/confocal-microscopes/confocal-microscope-with-multiplex-mode-lsm-980-airyscan-2.html#airyscan

## Slide 29
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Slow
Increasing speed of laser scanning confocal

## Slide 30
Scanners are slow
( single image can take ~1s)
Increasing speed of laser scanning confocal
Problem
Solutions
Resonant scanners ($$$, 30 frames/s or higher):
Mirrors that oscillate at a fixed (high) frequency
Less flexibility (rotation, zoom)
Uneven illumination issues
Increasing sensitivity allows higher speed
(with same signal-to-noise ratio)
Beware: speed costs photons (fewer from each location)
Scan less: trade field of view for speed
(cropping, linescans)
Scan coarsely: trade resolution for speed
(fewer pixels)
Notes: https://www.microscopyu.com/techniques/confocal/resonant-scanning-in-laser-confocal-microscopy

## Slide 31
(Less) sensitive spot detector
Fluorophores:
in focal plane
out of focal plane
Pinhole
Lenses, scanners et al
Laser
Damaging
Problems with laser scanning confocal

## Slide 32
Light damages cells and fluorophores
Decreasing damage of laser scanning confocal
Problem
Live cells (“photodamage”)
Fluorophores (“bleaching”)
Notes: https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3843244/

## Slide 33
Light damages cells and fluorophores
Decreasing damage of laser scanning confocal
Problem
Solutions
Improve sample:
Live: use brighter, more stable fluorophores
Fixed: use mounting media
Increasing sensitivity allows less damage
(with same signal-to-noise ratio)
Scan coarsely: trade resolution for less damage
Scan faster: trade sensitivity for less damage
Use longer wavelengths (less damaging)
Notes: https://www.microscopyu.com/techniques/confocal/resonant-scanning-in-laser-confocal-microscopy

## Slide 34
Pyramid of frustration
Contrast
Spatial resolution
Temporal resolution
Specimen health
?!

## Slide 35
Parameters we adjust on a laser scanning confocal
For contrast
Laser power
PMT Gain
Averaging
Scan Speed
For resolution
Zoom
Image size
Pixel size
Frame size

## Slide 36
Parameters we adjust on a laser scanning confocal
For contrast
Laser power
PMT Gain
Averaging
Scan Speed
For resolution
Zoom
Image size
Pixel size
Frame size

## Slide 37
Optimizing contrast on a laser scanning confocal
Laser power
PMT Gain
High contrast
High bleaching
Low contrast
Low bleaching

## Slide 38
Balancing gain and laser power
Gain: 700
Laser: 0.45%
Gain: 650
Laser: 0.9%
Gain: 600
Laser: 2%
Gain: 550
Laser: 4.5%
Gain: 500
Laser: 11.8%
Gain: 450
Laser: 55%
Notes: AVG signal levels in the entire field within 2% of each other between settings
Highest laser power not always best option: take into account bleaching and fluorophore saturation

## Slide 39
Optimizing contrast on a laser scanning confocal
Laser power
PMT Gain
Averaging
Scan Speed
High contrast
High bleaching
Low contrast
Low bleaching
High contrast
Low speed
Low contrast
High speed

## Slide 40
Averaging
No averaging
Average of
2 frames
Average of
4 frames
Average of
8 frames
Average of
16 frames
Notes: AVG signal levels in the entire field within 2% of each other between settings
Mention that speed has very similar effects

## Slide 41
Optimizing contrast on a laser scanning confocal
Laser power
PMT Gain
Averaging
Scan Speed
High contrast
High bleaching
Low contrast
Low bleaching
High contrast
Low speed
Low contrast
High speed
High(er) bleaching
Low(er) bleaching

## Slide 42
Beware of signal saturation
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 43
Saturation irreversibly destroys data
Distorts morphology
Compare details here

## Slide 44
Saturation irreversibly destroys data
Makes quantification imposible
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 45
To avoid saturation
Find bright object in brightest experimental condition
Adjust for maximum pixel intensities at half the full range

## Slide 46
Optimizing resolution on a laser scanning confocal
What level of biological detail do you need?
Beware:
More is not always the right answer

## Slide 47
You may not need high resolution…
Good enough for counting nuclei
Overkill

## Slide 48
Not enough detail
Better
On the other hand…
Notes: Images taken on LSM700 on BPAE slide. Zoom of 1 in both cases

## Slide 49
Optimizing resolution on a laser scanning confocal
What level of biological detail do you need?
If the answer is as much as possible…

## Slide 50
Optimizing resolution on a laser scanning confocal
Point objects are blurred
Key concept #1
Notes: PSF figure modified from Fig 2 http://zeiss-campus.magnet.fsu.edu/articles/superresolution/introduction.html

## Slide 51
Optimizing resolution on a laser scanning confocal
Minimal distance to distinguish 2 point objects: Resolution
Key concept #2
Resolution in XY
Resolution in Z
Rxy and Rz depend on fluorophore l and objective NA
Rxy
Rz
z
Notes: XY Modified from Fig 1 https://www.olympus-lifescience.com/en/microscope-resource/primer/techniques/confocal/resolutionintro/

## Slide 52
Optimizing resolution on a laser scanning confocal
Optimal sampling in XY and Z
Key concept #3
XY: pixel size
Optimal pixel size ~
Rxy
3
“Nyquist criterion”

## Slide 53
Optimizing resolution on a laser scanning confocal
Optimal sampling in XY and Z
Key concept #3
Z: spacing between slices
Optimal interval spacing ~
Rz
3
“Nyquist criterion”

## Slide 54
Optimizing resolution on a laser scanning confocal
Must set pixel size correctly !
How?

## Slide 55
Parameters we adjust on a laser scanning confocal
For contrast
Laser power
PMT Gain
Averaging
Scan Speed
For resolution
Zoom
Image size
Pixel size
Frame size

## Slide 56
Image size:
Pixel size:
Sample
Image
x
y
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Nx
Ny
xi
yi
Setting pixel size correctly

## Slide 57
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Setting pixel size correctly

## Slide 58
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Directly user-controlled
Setting pixel size correctly

## Slide 59
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Increase frame size
Smaller pixel size
Slower imaging
More bleaching
Setting pixel size correctly

## Slide 60
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Setting pixel size correctly

## Slide 61
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Increase zoom
Smaller area
x
y
xi
yi
More bleaching
Smaller pixel size
Setting pixel size correctly

## Slide 62
Recommended workflow
Determine pixel size you need.
Decision based on size and spacing of structures of interest
Determine image size you need.
Decision based on size and spacing of experimental units
(example: tissue structures cells, organelles)
Set the zoom to get the desired image size
Set the frame size to get the desired pixel size
Adjust other parameters to balance:
  contrast
  bleaching
  speed

## Slide 63
There is more to confocal than laser scanning
Contrast
Spatial resolution
Temporal resolution
Specimen health
Laser scanning confocal
Spinning disk confocal

## Slide 64
Spinning disk confocal microscope
Video: https://www.youtube.com/watch?v=eue5g6iSmPU
Notes: http://zeiss-campus.magnet.fsu.edu/articles/spinningdisk/introduction.html
https://www.youtube.com/watch?v=eue5g6iSmPU

## Slide 65
Spinning disk confocal microscope
Notes: http://zeiss-campus.magnet.fsu.edu/articles/spinningdisk/introduction.html

## Slide 66
Spinning disk confocal microscope
No free lunch: pinhole cross-talk reduces axial resolution
Pinhole array disk
Lenses:
objective et al
Sample
Thin sample
Thick sample
Notes: https://www.yokogawa.com/us/solutions/products-platforms/life-science/spinning-disk-confocal/csu-w1-confocal-scanner-unit/

## Slide 67
Spinning disk confocal microscope
No free lunch: pinhole cross-talk reduces axial resolution
Notes: http://zeiss-campus.magnet.fsu.edu/articles/spinningdisk/introduction.html

## Slide 68
Spinning disk confocal microscope
No free lunch: pinhole cross-talk reduces axial resolution
Pinhole array disk
Lenses:
objective et al
Sample
Thin sample
Thick sample
Thick sample
Larger distance between pinholes
Notes: https://www.yokogawa.com/us/solutions/products-platforms/life-science/spinning-disk-confocal/csu-w1-confocal-scanner-unit/

## Slide 69
Confocal is great !
But…

## Slide 70
Confocal is not the best technique for every application

## Slide 71
1mm
Nerve terminals in the cochlea
Fitzpatrick lab (Pappa et al, 2019)
50 mm
Widefield
Confocal
Fast
Sensitive
Great for thin samples (typically <10um)
Slow
Less sensitive
Great for thick samples (typically >10um)
Estimating 3D distance between DNA loci
Calabrese lab (Schertzer et al, 2019)
Notes: Calabrese lab paper - https://doi.org/10.1016/j.molcel.2019.05.028

## Slide 72
Thick things can be very hard to image
Why?
Light scattering due to heterogeneity in refractive index degrades transparency
Options?
Image thin things:
Cells cultured on surfaces
Slices of thicker things
Image transparent things
Naturally transparent
Chemically cleared
Use infrared light (“multiphoton”)

## Slide 73
Other issues when imaging

## Slide 74
What does “representative” image mean?
What population are we making inferences about?
Random sampling vs picking areas that “look right”.
Showing something can happen vs how often it happens.
Images as evidence vs images as illustrations.
“Publication quality” vs actual quality.

## Slide 75
How many images ?
5, 10, 20 ? Why?
Image more cells in more locations, more dishes, more cultures, more slices, more animals ? Where is the variability? More images within a dish may be a poor use of your time.
Do power analysis:
  Variance of what we are measuring
  Minimal effect we want to uncover
  Statistical test we are going to use
  Probability that we will find an effect if there is one

## Slide 76
Where can you get help ?
Imaging cores on campus:
Microscopy Services Laboratory (Rigor and Reproducibility section)
Neuroscience Microscopy Core
Hooker Imaging Core
Biology Department Microscopy Core
Online resources:
iBiology lectures on microscopy
Molecular expressions – exploring the world of optics and microscopy
Olympus microscopy resource center
Nikon MicroscopyU
Zeiss microscopy online campus
Good books ($):
Fundamentals of Light Microscopy and Electronic Imaging
Quantitative Imaging in Cell Biology
Great courses ($$$):
Analytical & Quantitative Light Microscopy (Woods Hole, MA)
Optical Microscopy & Imaging in the Biomedical Sciences (Woods Hole, MA)
Quantitative Imaging: From Acquisition to Analysis (Cold Spring Harbor, NY)
Quantitative Fluorescence Microscopy (Bar Harbor, ME)