## Slide 1: Confocal microscopy tips
Pablo Ariel
Microscopy Services Laboratory, UNC-CH
May 25th, 2017

## Slide 2
Before imaging
During imaging
After imaging

## Slide 3
Before imaging
During imaging
After imaging

## Slide 4
Optimize:
Antibody concentrations
Fixation, permeabilization
Incubation times, temperatures
Sample prep is critical

## Slide 5
Select proper fluorophores
AlexaFluor dyes are bright and photostable
AF488
AF546
FITC
Cy3
https://www.thermofisher.com/us/en/home/brands/molecular-probes/key-molecular-probes-products/alexa-fluor/alexa-fluor-dyes-excellent-photostability.html
Notes: Figure 1. Comparison of the photobleaching rates of the Alexa Fluor® 488 and Alexa Fluor® 546 dyes and the well-known fluorescein and Cy3 fluorophores. The cytoskeleton of bovine pulmonary artery endothelial cells (BPAEC) was labeled with (top series) Alexa Fluor® 488 phalloidin and mouse monoclonal anti–a-tubulin antibody in combination with Alexa Fluor® 546 goat anti–mouse IgG antibody or (bottom series) fluorescein phalloidin and the anti–a-tubulin antibody in combination with a commercially available Cy3 goat anti–mouse IgG antibody. The pseudocolored images were taken at 30-second intervals (0, 30, 90, and 210 seconds of exposure). The images were acquired with bandpass filter sets appropriate for fluorescein and rhodamine.

## Slide 6
Select compatible mounting media
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/cellular-imaging/fluorescence-microscopy-and-immunofluorescence-if/mounting-medium-antifades/prolong-gold-antifade.html
Notes: Photobleaching resistance for fluorescent dyes and proteins with ProLong antifade reagents. Photobleaching resistance was quantified on a LSM 710 (Zeiss) confocal microscope. HeLa or U2OS cells were stained and mounted using standard ICC protocols. Five regions within three fields of view were scanned 50x with a 1.58 μs dwell time per pixel. Excitation wavelength and intensity were optimized for the fluorophore being measured.

## Slide 7
Select compatible mounting media
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/cellular-imaging/fluorescence-microscopy-and-immunofluorescence-if/mounting-medium-antifades/prolong-gold-antifade.html
Notes: Photobleaching resistance for fluorescent dyes and proteins with ProLong antifade reagents. Photobleaching resistance was quantified on a LSM 710 (Zeiss) confocal microscope. HeLa or U2OS cells were stained and mounted using standard ICC protocols. Five regions within three fields of view were scanned 50x with a 1.58 μs dwell time per pixel. Excitation wavelength and intensity were optimized for the fluorophore being measured.

## Slide 8
Beware hardening mounting media
Images courtesy of Patrina Pellet (GE Healthcare)

## Slide 9
Use a #1.5 coverslip
#1
#1.5

## Slide 10
Use a #1.5 coverslip
http://olympus.magnet.fsu.edu/primer/anatomy/coverslipcorrection.html
Average difference between #1 and #1.5 coverslip

## Slide 11
Avoid large coverslips
Slide tilt

## Slide 12
Include controls
Examples:
No antibody
No primary
No target protein

## Slide 13
How will you sample?
What is “representative” ?
Random sampling?
Criteria for excluding cells?
Z stacks vs single planes?

## Slide 14
Before imaging
During imaging
After imaging

## Slide 15
Loading the proper configuration

## Slide 16
Loading the proper configuration
Smart Setup is not that smart

## Slide 17
Loading the proper configuration
Smart Setup:
DAPI + AF488 + AF594
“Best signal”
MSL_DAPI_AF488_AF594
(filters tweaked to match)
5 seconds
14 seconds
Notes: All settings are the same except for filters in the initial comparison
Filters modified to be the same in the second comparison.

## Slide 18
Selecting an objective

## Slide 19
Selecting an objective
Work backwards from the analysis you want to do
Size of objects of interest ?
Number of objects per field ?
Subcellular colocalization?
Dim?
Bleaches quickly?
Alive?
Very thick?
Consider your sample

## Slide 20
Size of objects of interest
Selecting an objective
Example 1: nuclei
10X/0.45
Zoom 1
40X/0.95
Zoom 1
Good enough for counting nuclei
Overkill
Notes: Images taken on LSM700 on BPAE slide. Zoom of 1 in both cases

## Slide 21
Size of objects of interest
Selecting an objective
Example 2: mitochondria
10X/0.45
Zoom 1
40X/0.95
Zoom 1
Not enough detail
Better
Notes: Images taken on LSM700 on BPAE slide. Zoom of 1 in both cases

## Slide 22
Selecting an objective
Resolution depends on NA, not Mag
Resolution depends on NA, not Mag

## Slide 23
Selecting an objective
40X/0.6
40X/1.3
Higher resolution means more detail
Resolution depends on NA, not Mag
Notes: Images taken on LSM700 on BPAE slide. 0.136um pixel size (roughly Nyquist). Conditions to roughly match signal and signal to noise levels in both images

## Slide 24
Number of objects per field
Selecting an objective
10X/0.45
20X/0.80
40X/0.95
Tradeoffs between number of cells, detail and time

## Slide 25
For subcellular colocalization
use Plan-Apochromat objectives
Selecting an objective
63X/1.4
40X/1.4
LSM 700 & 710
LSM 710

## Slide 26
Dim, photosensitive, live samples
Selecting an objective
Brightness
NA4
Mag2

Notes: Same settings for acquisition and display, except for:
pinhole size, which was adjusted to 1AU in both. Using the 40X pinhole size in the 63X led to a pinhole size smaller than 1AU and an even dimmer image.
Pixel size which I couldn’t make exactly identical: 0.122um in 40X, 0.120um in 63X
Note that pixel size is slightly below Nyquist

## Slide 27
Dim, photosensitive, live samples
Selecting an objective
40X/1.4 is the brightest objective
40X/1.4
63X/1.4
Notes: Same settings for acquisition and display, except for:
pinhole size, which was adjusted to 1AU in both. Using the 40X pinhole size in the 63X led to a pinhole size smaller than 1AU and an even dimmer image.
Pixel size which I couldn’t make exactly identical: 0.122um in 40X, 0.120um in 63X
Note that pixel size is slightly below Nyquist

## Slide 28
Objectives for special applications
Selecting an objective
25X/0.8
40X/0.6
LSM 710
LSM 700
Thick, aqueous samples
Plates

## Slide 29
There is a lot to consider
We can help you decide
Selecting an objective
Notes: Same settings for acquisition and display, except for:
pinhole size, which was adjusted to 1AU in both. Using the 40X pinhole size in the 63X led to a pinhole size smaller than 1AU and an even dimmer image.
Pixel size which I couldn’t make exactly identical: 0.122um in 40X, 0.120um in 63X
Note that pixel size is slightly below Nyquist

## Slide 30
Pinhole

## Slide 31
Pinhole
Set to 1 AU for your longest wavelength dye
Keep physical size same for all channels

## Slide 32
Setting the correct pixel size

## Slide 33
Setting the correct pixel size
Image size:
Pixel size:
Sample
Image
x
y
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Nx
Ny
xi
yi

## Slide 34
Setting the correct pixel size
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm

## Slide 35
Setting the correct pixel size
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Directly user-controlled

## Slide 36
Setting the correct pixel size
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Increase frame size
Smaller pixel size
Slower imaging
More bleaching

## Slide 37
Setting the correct pixel size
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm

## Slide 38
Setting the correct pixel size
Sample
Image
Zoom
Scan area of laser
Image size:
Pixel size:
X mm x Y mm
Frame size:
Nx pixels x Ny pixels
%
Xi mm x Yi mm
Increase zoom
Smaller area
x
y
xi
yi
More bleaching
Smaller pixel size

## Slide 39
Setting the correct pixel size
What is the correct size ?

## Slide 40
Setting the correct pixel size
Lower limit:
Pixel size ≈
Resolutionxy
2
≈
0.6 lem
2 NAobj
max

## Slide 41
Setting the correct pixel size
Lower limit:
Pixel size ≈
Resolutionxy
2
≈
0.6 520nm
2 1.4
Example:
40X/1.4
AlexaFluor 488
Pixel size ≈
220nm
2
≈
0.11mm

## Slide 42
Setting the correct pixel size
You may not need such small pixels…
Good enough for counting nuclei
Overkill

## Slide 43
Setting the correct pixel size
Beware of pixel sizes smaller than limit
Smallest recommended pixel size
Zoomed 2X smaller pixel size
Less area
More bleaching

## Slide 44
Less area
More bleaching
No extra detail!
Setting the correct pixel size
Beware of pixel sizes smaller than limit
Digital zoom of original image
Zoomed 2X smaller pixel size

## Slide 45
Setting gain, laser power & digital offset
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 46
Setting gain, laser power & digital offset
Do not saturate your image
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 47
Saturation irreversibly destroys data
Distorts morphology
Compare details here
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 48
Saturation irreversibly destroys data
Makes quantification imposible
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 49
To avoid saturation
Find bright object in brightest experimental condition
Adjust for maximum pixel intensities at half the full range
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 50
Setting gain, laser power & digital offset
What:
No saturated pixels
Digital offset
Gain vs laser power
How:
Range indicator
No pixels at 0
Good StoN
No bleaching
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 51
Setting gain and laser power

## Slide 52
Setting gain and laser power
Gain: 700
Laser: 0.45%
Gain: 650
Laser: 0.9%
Gain: 600
Laser: 2%
Gain: 550
Laser: 4.5%
Gain: 500
Laser: 11.8%
Gain: 450
Laser: 55%
Notes: AVG signal levels in the entire field within 2% of each other between settings
Highest laser power not always best option: take into account bleaching and fluorophore saturation

## Slide 53
Setting gain and laser power
Beware: high laser powers can cause bleaching
Test bleaching of your fluorophore in your sample
Note: laser powers not the same between LSM700 and LSM710
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 54
Averaging
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 55
Averaging
No averaging
Average of
2 frames
Average of
4 frames
Average of
8 frames
Average of
16 frames
Notes: AVG signal levels in the entire field within 2% of each other between settings
Mention that speed has very similar effects

## Slide 56
Bidirectional mode: 2X speed boost
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 57
Bidirectional mode: 2X speed boost
40X/1.3, zoom=2.3
Unidirectional
Bidirectional
Check for artifacts
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 58
Setting Z intervals
For maximal resolution use software suggestion
Size of objects of interest ?

## Slide 59
Before imaging
During imaging
After imaging

## Slide 60
Saving
Use .lsm or .czi
Do not use .tif
Do not export contents of image window

## Slide 61
Deconvolution improves contrast
Autoquant
Courtesy of Anastasia Zlatanou
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 62
Display adjustments
Make defensible adjustments

## Slide 63
Display adjustments
Adjustments MUST be same across images you will compare

## Slide 64
3D analysis in Imaris
Raw data, interpolated in 3D
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 65
3D analysis in Imaris
Raw data, interpolated in 3D
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 66
3D analysis in Imaris
Extract surfaces from data
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 67
3D analysis in Imaris
Extract colocalizing volumes from data
Notes: AVG signal levels in the entire field within 2% of each other between settings

## Slide 68
We are here to help YOU with YOUR science
Please ask questions (now or later)
msl@med.unc.edu