## Slide 1: Diffusion MRI - Lesson 3 DWI/DTI QC
Martin Styner UNC
Thanks to Guido Gerig, NYU
And many, many folks

## Slide 2: Content
Today’s topic: Diffusion MRI QC
  Still under-appreciated
Artifact correction with FSL:
  Motion & eddy current
  Susceptibility artifacts
  NRRD – NIFTI conversion
Remove bad DWI data with DTIPrep
Investigative tractography in Slicer

## Slide 3: DWI Artifacts
Diffusion images are sensitive to a number of artifacts
Eddy-current distortions
Susceptibility artifacts
Noise/SNR issues
Vibrational artifacts
Motion
Venetian blind artifacts
“unknown”…
DTIPrep: Remove bad DWI’s

## Slide 4: Bad DWI: Slicewise Artifact
Bad DWI detection => 3 choices
Remove bad DWI data (DTIPrep)
  Simple and clean, but removes partially ok data
DWI model correction
  RESTORE: DTI model, during estimation
  HARDIPrep (NYU/Gerig): Correction of corrupt slices (identified by DTIPrep) via spherical harmonic model
  FSL: Correction of corrupt drop-out slices (different identification) via Gaussian process model
  Both: Need motion & eddy current correction done
Don’t care…
  Estimated data significantly biased => incorrect results
  Berl et al (2015). Investigation of vibration-induced artifact in clinical diffusion-weighted imaging of pediatric subjects. Hum Brain Mapp, 36(12), 4745–57.

## Slide 5: Within DWI Motion Artifact
Venetian blind artifacts
DTIPrep: Detects and removes DWI with motion
  Differences between interleaved volumes
FSL: Detects motion between slices and correct them (only new version 5.12 or 6.1, not 6.0)
  Slice motion model as part of eddy current/motion correction, performs slice to volume registration
  Very slow
  High error setting for registration
  I would not use (at this time)

## Slide 6: Modern dMRI QC Workflow
DICOM conversion
(better) Remove bad DWIs via DTIPrep
Susceptibility artifact computation FSL
  Concatenation of corrected AP & PA
Brainmask
Motion & Eddy current correction
  Also applies susceptibility un-deformation
  Alternatively correct bad DWIs
(simpler) Remove bad DWIs via DTIPrep
DTI estimation – tractography QC

## Slide 7: DICOM convertion
Diffusion tags in DICOM are a mess!
  No field permeating standard yet
No single great converter tool
Currently Best: DWIConvert (Slicer)
  Captures many vendors and special cases well
  Can write both NRRD and NIFTI images
  Can convert to/from NRRD vs NIFTI
  Uses B-matrix for actual gradient and b-value (option)
Alternative: dcm2nix
  Converts to NIFTI, uses prescribed information

## Slide 8: NIFTI vs NRRD DWI
NIFTI
  4D – Image without diffusion information
  Gradients (b-vectors) & b-values in separate text files
  Confusing: often bvec & bval files not name linked
  No measurement frame
    b-vectors need to be specified in image coordinate space
    Vendors use different spaces (patient vs scanner coordinate space) => Adaption of bvecs needed

## Slide 9: NIFTI vs NRRD DWI
NRRD
  Single file format (.nrrd) or raw+header (.raw +.nhdr)
  All image information in header including diffusion data
  No confusions
  Measurement frame info
  Inconvenient: b-vector and b-value info combined into single value

## Slide 10: Susceptibility Correction
Correction of susceptibility artifacts
  Geometric distortion due to air/tissue interface or other sources (e.g. metal plates)
    Anatomy/Object induced distortion
  At first approximation: constant for all DWI
  Alternative: twice refocused sequence approach
AP
PA
Corr

## Slide 11: Susceptibility Correction
Needs two acquisitions with opposing polarities of the phase-encode direction
  AP vs PA, or LR vs RL
  Alternatively use structural T2w image or field map
FSL Topup
  Computes geometric un-distortion field
    Works for both opposing polarity or with structural T2w
  Time-consuming
  Application best with joint motion corr (only new FSL)
Not crucial for “Quick” QC
  Only used for minimal processing, not at QC
  ”Quick” QC separate for AP and PA

## Slide 12: Topup
Best: only use b=0 images (high SNR)
topup --imain=<B0s> --out=<fieldcoef> --fout=<field> --iout=<B0corr> --datain=<topupParam> --config=<fslregParam>
Needs 2 parameter files
  fslregParam: describes registration settings, defaults from FSL work well for most data
  topupParam: lists phase encoding & readout time
    Readout time is commonly consistent => bogus value ok
    Absolute sign irrelevant, relative sign is important
    E.g. 2 B0 AP & PA (orientation RAS) :
      0 1 0 0.1
      0 1 0 0.1
      0 -1 0 0.1
      0 -1 0 0.1

## Slide 13: Topup
Topup computes the un-deformation field
  Incorporate b0 image motion correction
  applytopup (old style) to apply to non b0
    Does NOT correct for motion, assumes no motion
    Avoid
  Use eddy_openmp (new style)
    Jointly estimates motion & eddy currents, as well as applies the susceptibility correction
    Compute on full concatenated data
    Very slow (use longleaf)

## Slide 14: Brain Masking in DWI
Options depending on availability
  Propagating sMRI brain mask to geometry corrected average b0 (affine reg)
  FSL bet to IDWI image
    Geometric average of non-b0 images
    Contains little non-brain tissue
Most will need manual correction
  Quick
  Check on FA for non-inclusion of skull
  I suggest to use ITK-Snap

## Slide 15: DWI Motion Correction
Quantifies and corrects motion between DWIs
  Both DTIPrep and FSL
  Reporting of detected/corrected motion
  Analysis of motion presence wrt clinical variables necessary nowadays
Old FSL (pre 5.12 and 6.0)
  Registration to first b0
DTIPrep
  Registration to unbiased b0 average
  Estimate motion better for later acquired DWIs
  Better than old FSL (Kreilkamp 2015 JMRI)

## Slide 16: Eddy Current Artifacts
Different for every DWI
  Shearing geometric distortion
No/little artifacts for b0 images
First order approximation
  Correct via affine registration
  Incorporates motion correction
  Nearly all tools do it this way

## Slide 17: New FSL Eddy
Only in 5.12/5.13 and 6.1 (and higher)
eddy_openmp does everything
  Eddy current correction
  Motion correction
  Susceptibility field application
  Within DWI motion correction (do not use)
  Within DWI drop out correction (prefer removal)
Needs brain mask

## Slide 18: New FSL Eddy
eddy_openmp
  Very slow (> 1 day on single core)
  Can use multiple cores (do so)
    Environment variable: OMP_NUM_THREADS 4
    bvals, bvecs, topup para, “--data_is_shelled”
  Needs new configuration file identifying registration target per DWI by index (base 1)
    Simple: all 1 => old FSL, register to first b0
    Better: index of last acquired b0
      Example: 5 b0, equally spaced, 35 total
        1 1 1 1 1 1 1 2 2 2 2 2 2 2 3 3 3 3 3 3 3 4 4 4 4 4 4 4 5 5 5 5 5 5 5

## Slide 19: DTIPrep
DTIPrep – automatic diffusion QC tool
  Protocol based QC and correction
Image dimensions check
Gradients directions check
Intensity changes across slices
Excessive motion across slices
Motion and eddy current correction
Residual motion detection
Vibration/direction DTI artifact detection
  GUI and command line
  Summary reporting, including motion

## Slide 20: DWI info
Detailed
DWI
Info

## Slide 21: Gradient info in 3D
Displays gradient scheme on unit sphere
(F = File)
Check for uniformity

## Slide 22: Visual DTI QC
Visual in 3D Slicer:
  Remove DWI’s manually with significant residual motion
    Happens rarely
  Reconstruct tensor and perform investigative tractography
    Ensure appropriate tractography
    Major tracts, spatially diverse
    CC genu, CC splenium, CC tapetum, uncinate, fornix, inferior longitudinal, cerebro-spinal, anterior cingulate, (arcuate)

## Slide 23: Conclusion
Diffusion QC is a must
  Otherwise you get crap data
  FSL finally has detection of bad DWI!
    Hurray!
  FSL does not remove, but correct bad DWI and within-DWI motion
    Ehhh, that’s not what I wanted…
  Best: Combo of FSL/HARDIPrep and DTIPrep
DWI QC is now very very time consuming

## Slide 24: Slicer Demo